# Program GuiDe	Description (Zimm Plots)

Why are you here?  You are probably here to learn about Zimm plots.  GuiDe is an alternative to doing the classical Zimm analysis, but this document does contain a nice section on Zimm plot theory, which begins on Page 17.  (If you are reading this in Word, hit Function Key 5, then type 17).

Technical Description.  GuiDe is a Visual Basic program for Windows 95.  GuiDe reads .GAL files (“galvanometer” files, although they may contain count rate information, not photocurrent) prepared by SSLSC.BAS.  Dr. Stryjewski’s Phoenix program for the multidetector device will be compatible also.  Results may be printed to paper or stored in a file.  GuiDe can prepare .DAT (data) files for the publication quality graphics program, Origin, from Microcal.  These files can also be used to treat the data in ways not supported by GuiDe.  GuiDe can also produce .TXT files which contain a summary of the results and can be read by a simple word processor, such as Notepad.

Purpose.  GuiDe is for analysis of “data grids” collected during a static light scattering (SLS) experiment.  There are two independent variables in such experiments, concentration, c, and scattering angle, The latter variable is related to the scattering vector magnitude, q.  The dependent variable is scattered intensity.  All this information, plus ancillary data (wavelength, temperature, time, date, etc.) come to GuiDe in a .GAL file.  From these, GuiDe extracts the polymer molecular weight, M, second virial coefficient, A2, and radius of gyration, Rg.  A very important function of GuiDe is to eliminate “off-trend” data, which it does with a convenient graphical interface.  There are lots of other things one could do the data grid, and GuiDe facilitates unconventional applications by converting the raw data to .DAT files in column form for Origin or a spreadsheet or other plotting program.

Relation to Other Programs.  Our lab has used these other programs for analysis of SLS data:

- Zimm.PAS, a Turbo Pascal program developed in the mid-1980’s by Paul Russo.  Zimm.PAS only has low-resolution graphics and is hardly user friendly (though it has an appealing logic and efficiency once you are used to it).  The original purpose was simple plotting of the data grid, which was intended to be manually analyzed (with rulers or French curves).  Mostly for the heck of it, a linear fit of the data in the grid was added.  Of course, users instantly relied on this--with rather bad consequences for large random coils, which produce nonlinear Zimm plots.  Within the limitations of its linear fit, Zimm.PAS does a good job of handling uncertainty in the data.  It is relatively bug-free.  There is no data editing; this must be accomplished with a word processor.  Zimm.PAS was never updated because newer releases of Turbo Pascal were not backward compatible, requiring a large effort.

- Zimm.VB, a Visual Basic program developed about 1994 by Drew Poche’, then at Dow.  It can perform polynomial fits to the data grid, and is much more appealing and user friendly than Zimm.PAS.  The program is a very good example of conventional Zimm analysis, for it has a lot of nice features--like overlay of data, a good polynomial fit routine and intuitive user interface.

- DAWN and ASTRA.  These programs, supplied by Wyatt for the DAWN apparatus, are pretty slick in their latest incarnations.  They can perform a Zimm-style analysis of the data grid, and they do this with three methods:  Zimm, Debye or Berry.  It is important to understand that the Debye method in these programs is not the same concept as the Debye method in the GuiDe program.  Similarly, while there is a “Zimm plot” in GuiDe, it is important to realize that this is for display purposes only, and not functional.

Explanation of General Approach.  GuiDe represents an unconventional approach to SLS data.  All the programs mentioned above, and most others in the light scattering community, follow the lead of Zimm’s famous paper.  (Zimm, B.H. The Scattering of Light and the Radial Distribution Function of High Polymer Solutions. J.Chem.Phys. 16(12):1093-1099, 1948; see also Zimm, B.H. Apparatus and Methods for Measurement and Interpretation of the Angular Variation of Light Scattering; Preliminary Results on Polystyrene Solutions. J.Chem.Phys. 16(12):1099-1116, 1948.)  Zimm’s brilliant idea is explained in Appendix 1, Theory.  In brief, Zimm’s solution to graphing a function of two independent variables was to sum the two independent variables, taking care to scale one of them so that each contributes significantly to the abscissa.  Thus, the quotient Kc/( (see Appendix 1 for these very conventional terms) was plotted on the y-axis against an x-axis of sin2(2) + kc.  The scaling factor k (units of inverse concentration) was selected to give the concentration variable an importance comparable to the sin2(2) term, which ranges from 0-1.  This results in a tilted data grid, a single plot that facilitates the c=0 and =0 extrapolations.

One may speculate as to why the Zimm plot became such a fixture in polymer science.  The necessary extrapolations to zero angle and concentration could just as well have been performed in separate plots, as we shall see.  These plots would be easier to construct--even in Zimm’s day of manual plotting--than the single Zimm plot.  They would require more space in journals than a single plot, and journals of that era were wonders of economy compared to today’s bloated rags.  Whatever the reason, one cannot deny the appeal of difficult-to-obtain data aligned beautifully on a single grid.  Even dynamic light scattering data, which also depends on angle and concentration, have been plotted in the style of Zimm.  The GuiDe program still shows data in the traditional Zimm form, but GuiDe does not begin with a Zimm plot of the data grid, nor does it rely upon it to obtain the answers.

Instead, there appear two separate plots--one showing c/( vs. c and the other c/( vs. sin2(/2) (c/( is plotted and not Kc/( because often one wishes to evaluate the data quality before the optical constant K is known--i.e., before the dn/dc measurements have been performed).  Whole concentrations or whole angles can be deleted with a few mouse clicks.  Individual points can also be deleted.  The “surviving” points can be viewed.  It is easier to do all this on a conventional two-dimensional plot than it is on a Zimm plot tilting at some arbitrary angle, depending on the scaling constant.  Color-coded data points also help.

What happens next is also unconventional.  For each concentration, the angular dependence of ( is fit to functions derived by Debye or Guinier, or both.  This results in the Rayleigh factor scattered to zero angle and the apparent radius of gyration.  These two parameters are plotted against concentration.  The intercept of the Rg vs. c plot provides the true Rg.  The intercept and slope of the ((q=0) vs. c plot provide the molecular weight and virial coefficient, respectively.  Doing things in reverse (fitting towards zero concentration first, then plotting the results against angle) is not supported.

Finally a Zimm plot is constructed.  The data are plotted and a variable scaling factor applied to adjust the tilt of the grid.  Then follows the overlay of the computed lines (even though polynomial or non-linear functions may have been used).  Guidelines can be drawn and the file printed out to simplify good, old fashioned graphical analysis.

All this may seem like re-inventing the wheel.  The first part of GuiDe is, indeed, just convenience:  identifying outlier data points is simply easier with orthogonal axes.  The more substantive part of the program is that it facilitates using the Guinier plot and the Debye random coil function.  The Guinier plot is an excellent device when the shape is unknown.  Rg is obtained from the slope of a semi-log plot (ln(() vs. q2, where q=scattering vector magnitude) without propagation of error from the intercept.  When data are not available at sufficiently low q, the Guinier plot may curve.  GuiDe makes it easy to select the linear region and/or take the initial slope of some polynomial fit.  The Debye coil form is useful in the common case that the sample is known to have a random coil configuration and the data were not obtained at sufficiently low values of qRg to produce a linear Guinier plot.  This optimizes the use of the available data.  The square root plots invented by Berry are not supported in GuiDe:  with a general polynomial fit available, the need to obtain linearized graphs is reduced.  Like the Zimm plot, the Berry plot is designed to linearize data.  There is no need to do this anymore.

Limitations.  GuiDe’s limitations are many.  It is completely unsophisticated about handling random uncertainties in data.  However, systematic uncertainties are usually greater anyway.  There is no reason why we could not extrapolate to zero c first (rather than zero q as we do now).  There are undoubtedly “bugs” in the program, too.  Please report errors and make suggestions so it can be improved.

Installation.  A GuiDe distribution diskette may be obtained from Paul Russo.  Insert it and type Setup or click the Setup icon.  This will install GuiDe.EXE in a directory C:\GuiDe, along with several sample .GAL files (with special emphasis on poor ones).

Running GuiDe with the Sample Files.  Double click the GuiDe icon to start the program.  The screen will look like this:

Program Design.  This is a good place to point out some of GuiDe’s features.  Writing a program for repetitive scientific analysis is not like writing a word processor.  Often, there is a single, correct way to proceed.  Old-fashioned sequential programming (Fortran, Pascal and QuickBasic in order of disappearance) is very good at this, but also unfriendly.  In contrast, Visual Basic is (almost) an object oriented language, and the user is bombarded with menus, buttons and so forth.  It is entirely possible to do things out of order in such a program.  GuiDe attempts to keep the user “on the straight and narrow” by relying mostly on Command Buttons that are enabled or disabled as appropriate.  There are not many hidden options or drop-down menus.  Many screens have a light purple area, in which instructions are written.  If things get done out of order, GuiDe attempts to correct for it, but the user should not push his/her luck!  Try to follow the instructions in the purple boxes.  Do not be afraid to restart the program if you are worried about its behavior.

Open a .GAL file.  Note that many of the buttons are not yet “enabled” (you can read their gray captions, but clicking on them has no effect).  The Open .GAL button is enabled, and so is the Exit button.

- Click the Open .GAL button.  Most buttons can be accessed by key chording, using the ALT key plus the letter that is underlined on the button.  Thus, you could type ALT-O instead of clicking with the mouse.  A dialog box will open up and, if you are used to using Windows, it will be clear what to do:  just browse to find the .GAL file you wish to open.  Double click it or use OK to open the file.

For now, double click on the PSBAD.GAL file.  A message box appears describing this sample briefly--how many angles and concentrations, etc.  Clear this box by clicking OK.

Data Selection.  Note that other buttons become “enabled” (their captions turn from gray to black) as the selected file is plotted.  The screen will look like this:

This is a bad data set (concentration errors).  You can edit data in 3 ways:

- Remove a whole concentration.  Click the “Kill Concentration” button.  Then push the “up” arrow on the blue “spin bar”.  Note that a whole set of concentrations turns red in color.  If you wish to kill this set (for example, the “first” concentration is not the lowest, which probably means an error in sample preparation) then hit the “Yes, Kill” button.  The points acquire a yellow center.

- Remove a whole angle.  Click the “Kill Angle” button.  Proceed as above.

- Remove a single point.  Click on it, and a dialog box will appear asking if you are sure you wish to remove the point.

Do not worry much about the smallness of the screen; you will soon see the data in excruciating detail.  The size is adequate for the purpose of data selection:  any data not clearly goofy shouldn’t be deleted anyway.

You can restore data by hitting the “Revert” button.  This restores all the deleted data.

You can focus in on just the retained points by clicking “See Survivors”

Note.  The method by which GuiDe “kills” bad points is to turn the sign on the intensity (not the Rayleigh factor) negative.  Thus, the sign of the intensity is used throughout the program code to determine whether a point is plotted, fitted, etc.

Analyzing Data by the Guinier Method

1.  Set the points.

- Start by clicking the “Guinier” button.  The data will be replotted on a new, larger form as ln(() vs. q2.  It looks like this:

In this form, all concentrations are shown at once.  You can click the arrows on the “spin bars” to control the scales of the plot.

- Now click the “Analyze” button.  Note that its caption turns to “Next Concentration” as the first concentration Guinier plot appears, magnified.

- Your objective is to find the linear regime of this plot.  Perhaps all of it is linear, or perhaps not.  Or maybe there is no trend at all, indicating a small polymer.  Click on the first point you wish to include and then set Lower q Limit.  Click on the last point you wish to include and then click Upper q Limit.  The data are replotted in green to indicate linear regime  points and red to indicate points outside the linear regime.

- Adjust the order of the fit (default is linear, and this is the usual intent of the program:  find the linear regime and use it).

- Note that the Rg value for this concentration appears in a box to the right of the plot.

- Then click “Next Concentration” and repeat.

- Eventually, a dialog box will tell you that all concentrations have been set.  At this point, you have established the key parameters of your data grid--((q=0) and apparent Rg for each concentration.

- If you like what you did, go ahead and click “Continue”.  Otherwise hit “Return” to go back to the main menu, click “Guinier” and try again.

2.  Obtain the True Radius of Gyration

- The apparent Rg is plotted against concentration, like this:

- Your plot may be linear or curved, and can be quite noisy.  A spinbar/textbox combination on the bottom of the curve controls the order of the polynomial fit, and the true Rg (Rg,app at c=0) from the intercept appears in another text box.  Do what seems reasonble....be conservative, because the data may be noisy, especially for small polymers.

- Click the “Forward....” button to deal with the zero-angle data, which yields Mw and A2

3.  Obtain ((q=0, c=0)

The plot looks like this:

- Adjust the polynomial fit order reasonably to obtain the intercept (inversely related to molecular weight) and slope (related to the virial coefficient).

- Click the “Compute Mw” button to proceed to the summary.

Summarize the Data

The form looks like this:

- Just fill in the dn/dc text box and Mw and A2 will be computed.

- If you don’t know dn/dc, you can look it up in Huglin (if you haven’t met Huglin yet, it’s time!) or elsewhere--but the wavelength will often be wrong, because lots of good measurements were made before the laser was commonly used.  So....you can click the “Compute” button and a new form will appear, like this:

- Type in up to 5 pairs of wavelengths () and dn/dc from the literature.  Click the “Plot” button, and a Cauchy plot (dn/dc vs. appears.  It will almost certainly be linear.  Type in the wavelength at which you desire to know (the form is pre-loaded with the wavelength stated in the .GAL file) dn/dc and hit “Plot” followed by “Compute”.  The computed point will appear on the plot, and the value in the appropriate text box.  The value also appears back on the “Summary” form.

- Hit “Return to Summary” and the results should already be there waiting for you.

Zimm Plot

As already described, the Zimm plot in this program is not used for analysis.  Rather, it shows how the three parameters obtained by other means can fit the data, when the data are in the traditional Zimm representation.  Moreover, it does this in a simple, linear way.  Even if you used higher-order polynomials (Guinier method) or the Debye P() form, the Zimm plot in this program only shows a linear representation according to the simplest, linear model

Kc/( = Mw-1(1 + q2Rg2/3) + 2A2c.

You can be the judge of how well this much simplicity works.

- First, adjust the scaling constant “k” up or down to make the grid look nice.  No cheating--avoid very large or very small values that make one set of lines too vertical.

- Then hit “Next Concentration” repeatedly to add these lines.  When you are done, the “Next Angle” button becomes enabled.  Hit this repeatedly to add the lines for each angle.  The fat, light blue circles show the points at q=0 or c=0, as obtained by the linear Zimm model (see equation just above).

- If the model lines don’t really hit the data grid, remember the limitations of what we are doing:  we aren’t trying to fit individual lines on the grid!  This is a GuiDe analysis.  So click the Help button for words of encouragement.  However, if there are serious mismatches between the data grid and the model, consider the possibility that the data are just poor.

- You can erase the plot and try again, if you wish.  Almost any Zimm plot can be made to look OK with the right scaling parameter.  But using too-large or too-small values is “cheating.”

- Finally, there is one very primitive option worth trying.  Click “Wipe Clean” to clear and replot the data grid.  Rescale it to look good.  Then click one of the Highlight Extrapolated option dots (often called Radio Buttons) located to the left of the y-axis.  The program will draw lines that will help you do a conventional, old-fashioned, graphical Zimm analysis.  The plot in this form could be dropped into another document and printed.  To do this, hit the PrtScr (print screen) key, open a new document with Word or other word processor, and hit “Paste” (ctrl-V).  Then click the Print button.  For mediocre data, such manual drawing is probably good enough....and a fine experience.

The following Zimm plot shows the lines from the linear Zimm model, Guinier method, plus the guidelines if you wanted to try to draw something by hand over the model lines.  It would ordinarily be better to skip the model lines when drawing by hand.  This is really lousy data, so the model doesn’t match it very well:  the model “smoothes” over the whole grid, whereas individual concentrations are very wrong in this data set.

Analyzing the Data by the Debye Method

- If you are not at the Main Menu/Data Selection form (where the program began) return there by pressing the appropriate buttons (usually in the upper right corner).  Press “Debye” to initiate the process.  As with the Guinier method, you first see data for all the concentrations at once.  You can scale the plot if desired.

- Press “Analyze” and note that, as a large plot appears for the first concentration, the caption of the “Analyze” button changes to “Next Concentration”.  This action recalls  the Guinier method.  However, the goal here is different:  you are not trying to find the linear regime.  Rather, you are trying to fit all the available data to a form appropriate for random coils:  P() = ( ()/((0) = (2/x2)[e-x -1 + x] where x = qRg.  This equation is nonlinear, and GuiDe executes a very simple two-dimensional grid search algorithm to find the parameters (((q=0) and Rg) that fit the data best.  By default, trial Rg values run from 50Å to 2000Å by default.  Trial values of ((q=0) vary from 50% below to 50% above the value ((q1) where q1 is the lowest-measured q value.  Only for really huge polymers measured at too-high q would this not work.

- The grid fit algorithm is a bit slow (say, 30 s for each concentration).  To speed it along, you can click the right mouse button and a pink form will appear:

- Set the parameters of the form sensibly for your polymer.  For example, by looking at the plots so far, you will know if it is a very small polymer.  In that case, there is no reason to search all of the high-Rg parameter space.  Instead, set the limit at something sensible, like 500 Å.  You can also adjust the “fineness” of the intensity fit, but this won’t save any time.

- Observe the quality of the fit; the plots look like this:

- For very small polymers, where the ((q) vs. q2 trend is not pronounced, you will observe that the low-q part of the fitted curve descends rather steeply compared to the rest of the data.  This may be cause for some suspicion; compare the Rg values to those obtained from the Guinier plot and consider whether you really should be using the Debye method at all.

- Perform the Summary and ZimmPlot steps as described above.

Publication Quality Graphics.  This is not really a function or concern of GuiDe.  However, the program can simplify the process of making the one or two beautiful plots you might need for a journal publication by writing out data from .GAL files in a form acceptable to the plotting/analysis program, Origin.

- Press “Return” to get back to the Main Menu/Data Selection menu (where the program began).

- Press the “Origin” button and follow the instructions to make two .DAT files--one with the Rayleigh factor tabulated as a function of angle for various concentrations and the other with Rayleigh factor tabulated as a function of concentration for various angles.  These plots are easily manipulated in Origin or another spreadsheet.

REMEMBER--The data used here to illustrate GuiDe really are terrible.  There is a video tape of a GuiDe session where the full editing power of the program is applied to this same data set to come up with more respectable behavior.  But don’t be misled by the videotape.  Light scattering is definiately a “garbage in, garbage out” experiment!

Appendix 1.  Theory.

Here we reproduce the information that used to go with the old Zimm.PAS program, somewhat modernized.

Traditional Zimm Approach

A static light scattering measurement on well-dispersed polymers in solution will yield molecular weight, M, radius of gyration, Rg, and virial coefficient, A2.  The precision of static light scattering is much lower than that of dynamic light scattering.  Additionally, size information is only obtained for particles larer than about 100 Å.  However, all parameters measured from static light scattering are absolute.  In dynamic light scatering, only the diffusion coefficient is absolute; size and/or molecular weight require assumptions.  A reasonably detailed and fundamental description of static light scattering is given in Chem 4595.

A Zimm plot consists of Kc/( plotted against q2 + kc, where K is an “optical constant”, c is the polymer concentration (as g/mL), ( is the Rayleigh factor, which is related to scattered intensity as described below.  The scattering vector magnitude, q, is given by

q = 4n sin(/2)/ 	{1}

where n is the solution refractive index, nearly the same as the solvent refractive index for sufficiently dilute solutions,  is the scattering angle and is the in vacuo wavelength.  The parameter k is an arbitrary scaling constant selected to give the concentration independent variable an importance comparable to that of the squared scattering vector.  Historically, lots of Zimm plots were plotted as Kc/( vs. sin2(/2) + kc.  The reason is, in the old days, you would have had to compute all those q2 numbers manually.  So instead, you just lived with sin2(/2).  Using q2 instead of sin2(/2) in the following equations makes them more compact and readable.

The basis of the Zimm plot is found in the fundamental expression of static light scattering in the limit of “not too large” particles (Rayleigh-Gans-Debye limit):

Kc/( = Mw-1(1 + q2Rg2/3) + 2A2c	{2}

Note that this form is a simplification of the full expression of Zimm.  The optical constant, K, is:

K = 42n2(dn/dc)2/o4Na	{3}

where Na is Avogadro’s number.  The factor 4 is for vertically polarized incident light.  For an unpolarized incident beam, replace it by 2(1+cos2) which is slightly inconvenient because then K = K().  The all-important parameter for static light scattering is the specific refractive index increment, dn/dc, which must be measured separately.  If dn/dc = 0, then there is no scattered intensity.

The dependent variable, Kc/(, depends on both c and q.  One may think in terms of a 3-dimensional plot:

Zimm devised a two-dimensional representation in which the two dependent variables, c and q2 are summed. To ensure that the q2 variable and the c variable contribute significantly, one of them (usually c) is scaled with a factor k.  In sum, the Zimm approach is to plot Kc/( vs. q2 + kc.  The result is a tilted grid of points:

In this hypothetical Zimm plot, there are 4 polymer concentrations and 6 different scattering angles.  (The different concentrations are represented by different colors, if you are looking at this document in color.)  By changing the value of the scaling constant k, it is possible to make this plot “fold in” on itself.  You have to be very careful when assessing the quality of a Zimm plot.  It’s not just a matter of appearance:  it is possible to hide data of poor quality by selecting low k values and, especially, very high ones.

The purpose of the Zimm plot is to assess the scattering behavior in two important limits.  1)  At c=0, the indepdent variable is just q2 and Eq. 2 reads:

Kc/( = Mw-1(1 + q2Rg2/3)	{4}

Thus, if the behavior at c=0 can be deduced somehow, we could determine molecular weight (inversely, from the intercept) and radius of gyration (from the slope and intercept).

Intercept(at c=0) = Mw-1	{5a}

(3 ( Slope/Intercept)1/2 = Rg	{5b}

The behavior at c=0 can be deduced simply by extending vertical guidelines up from the x-axis for each of the 6 q2 values.

Then lines are drawn through points at constant angle and different concentrations until it extends back to the appropriate guide line. Mark the intersection with a large cirlce (light blue if you are seeing this in color).  A line (or curve, if necessary) is drawn through these points to describe the behavior at c=0.  Mw and Rg are obtained from this curve, using Eq. 5a and 5b.

2)  At q=0,  Eq. 3 reads:

Kc/( = Mw-1 + 2A2c	{6}

So, from the intercept, one obtains the molecular weight (again) and from the slope, one gets the virial coefficient.

Intercept = Mw-1 	{7a}

k ( Slope/2 = A2		{7b}

The scaling constant k appears in Eq. 7b because the slope of a scaled Zimm plot at q = 0 represents Kc/( vs kc not vs. c.  Again, guide lines are used to extrapolate the behavior at q = 0 from the data grid.  Draw vertical guide lines at x = kc for the 4 concentrations.  Then draw lines (or curves, as needed) through grid points at constant concentration to obtain the Kc/( curve for q=0.  Again, draw big circles where the lines (curves) hit the guide lines.  Draw a line (curve) through this to obtain Mw-1 and A2.

Mechanics of Data Gathering and Automation

In this section, we consider the gathering of data (program SSLSC) and processing of data (programs Zimm.PAS, GuiDe.VB and Zimm.VB).  We fully develop the propagation of errors due to random sources (noise).  However, systematic errors (concentration errors, stray light, dust) are often more significant in static light scattering, so the noise treatment here is to some extent an exercise in error propagation mathematics and possibly not very useful, except in the limit where great care has been taken to eliminate systematic errors.

It is often more convenient to plot c/( rather than Kc/(, because the optical constant is not necessarily known at first.  All results are just multiplied by K later.  Our analysis is based on this approach.  The first objective is to obtain c/(.  c is easy enough; what about (?  To begin with, ( is very hard to measure directly.  Literally years of effort could be involved.  Most instruments built today are not well suited to the task, and it requires very precise physical measurements.  This is because the Rayleigh factor is:

( = Ir2/VIo	{8}

where I is the scattered intensity at some angle when the sample is illuminated by a beam of intensity Io.  V is the volume of solution detected, and the hard part to measure.  The other hard part to get right is the ratio I/Io, since Io is typically 105-106 time larger than I.  No photomultiplier tube is linear over this range, so very careful “splicing” of the phototube response function is required.  Little wonder that Rayleigh factor measurements were hotly debated from the inception of photometric measurements (in the 1940’s) until about the late 60’s.  Obviously, we would not care to measure Rayleigh factors for all our solutions and at all our angles.  The answer is to calibrate the light scattering instrument at one angle, using one of the known standards that earlier scatterbrains worked so hard to measure.  A suitable standard is toluene.  However, before we proceed, we need to say something about the optical layout we typically use for light scattering measurements.

Below appears an x-y-z coordinate system.  We usually send the laser beam in along the y-axis, with its electric field vertically polarized along the z-axis.  The detector arm swings around the horizontal x-y- plane.  Now, light of all polarizations is usually detected.  This arrangement is called the Uv geometry--i.e., unpolarized detection/vertical incidence.

The appropriate Rayleigh factor for toluene at 90o scattering angle, 23oC, is:

(uv = 	14.02 x 10-6 cm-1 @ 6328 Å (Helium-Neon red)

28.07 x 10-6 cm-1 @ 5320 Å (Frequency doubled diode)

32.08 x 10-6 cm-1 @ 5145 Å (Argon ion green)		{9}

39.64 x 10-6 cm-1 @ 4880 Å (Argon ion blue)

(Exercise:  how are these values related?)

Once this is known, it is possible to convert measured intensities to Rayleigh factors if the intensity of the Rayleigh standard is also known.  Essentiallly, it is done by a simple proportionality:

(measured / Imeasured = (standard / Istandard	{10}

However, certain correction terms are required, and the more complicated equation which calculates ( is:

(()=[I() - Is()](sin[(std(90)/Istd(90)]([Vstd/Vsample]	{11}

where it is assumed that the Rayleigh standard is measured at 90o.  All symbols I correspond to scattered light (dark current or dark count from the photomultiplier tube needs to be subtracted--not from the lead term where subtraction occurs automatically, but certainly from the denominator in [(std(90)/Istd(90)]).  The unsubscripted I means a solution intensity; Is refers to the solvent intensity.  The term sin adjusts for the angle dependence of scattering volume; a larger volume is seen at lower angles, so we must refer each measurement back to what we would see if we had the same volume as we would at 90o.  More generally, one would use a geometric factor G(), but our machine follows sin very closely.  The term (std(90)/Istd(90) converts intensities to Rayleigh factors.  The last factor (Vstd/Vsample) corrects for the fact that the volume detected depends not only on angle, but on the fluid itself!  The reason can be deduced by ray-tracing to determine what the detector sees.  It turns out that Snell’s law refractive index corrections soon appear, with the result that (Vstd/Vsample) = (nstd/nsample)m where m is 1 or 2 depending on the type of receiving optics.  If the detector sees beyond the beam substantially (i.e., the beam looks like a thin pencil of light to the detecor) then m=1.  If the detector looks well within the beam (i.e., a “fat” beam compared to detector size) then m=2.  In the first case, m=1 represents that refractive index corrections affect only the length along the beam that is detected:  the whole beam height is still seen.  In the second case, m=2 implies that both the height and length of beam detected vary with reractive index.  In most old instrumeents, m=2.  however, in our instrument, which also serves a dynamic light scattering device, the beam thickness can be very small (ca. 100).  Depending on the aperture/pinhole settings, one might have m=2 or m=2.  It is simple to tell just by looking into the ocular with both aperture and pinholes set as desied for measurement.  Note:  usually, a static light scattering experiment uses larger pinholes and apertures in order to raise the number of coherence areas detected, thereby deliberately destroying the coherent intensity fluctuations which are essential for dynamic light scattering but an obvious nuissance in a static measurement).  The following diagram shows m=1 and m=2 situations.

With the substitution for (Vstd / Vsample) we can easily see that the ordinate we wish to plot, c/(, has the form:

c/(() = 	{12}

Now, we should really also calculate the uncertainty in the ratio c/(.  We assume that c, , Rstd and nsamp are all known perfectly well.  (In other words, errors here fall into the systematic category, along with stray light, misalignment and concentration errors).  This is obviously not really true, and these systematic uncertainties actually may dominate the random intensity variations.  As stated already, this whole section is somewhat of an exercise in error propagation.  Anyway, under our assumptions, c/( is a function of I(q), Is() and Istd.

c/( = f(I(), Is(),Istd)		{13}

By standard propagation of error formulae (see, for example, the classic text “Data Reduction and Analysis for the Physical Sciences” by P. R. Bevington) the mean square uncertainty, , is given by:





These values can be used in the least square fits to constant  and constant c lines.  Each fitted line has the form y=Mx+B and M and B are computed.  (M should not be confused with the molecular weight; it is just a slope; neither should M be confused with the volume exponent, m).  The purpose of these lines is to extrapolate to a given abscissa value x’; for example, x’ can be q2 or kc, depending on whether the guideline is for c=0 or q=0, respectively.  Suppose we want the c=0 extrapolation for data taken at q2 corresponding to 45o scattering angle.  Then x’ = q245 and the extrapolated c/( value is:  c/((c=0, q45) = M45x’ + B45.  Now the initial temptation is to compute the uncertainty of this extrapolated point as:



However, this equation overestimates c/( because the intercept is extrapolated all the way back not to x’ but to 0.  Then the uncertainty is projected back all the way to x’ when, in fact, the data shift of the Zimm plot itself is exact:  exactly q245 in the example here.  Thus, due to the arbitrary (and exact) data shift, the intercept is being obtained by an unrealistically long extrapolation, and is therefore made less precise.  There is also the problem (possibly smaller) that the uncertainties in M and B are not strictly independent but covariant, in which case simple propagation of error formula like that above is invalid.  The solution is to “undo” the arbitrary shift before a given line at constant angle (or contant concentration) is fitted.  This gives very relistic uncertainties for c/(c=0,q=const or c/(q=0,c=const directly as the intercept uncertainties. This is what Zimm.PAS does.  The failure of Eq. 16 was immediately obvious in an old version of Zimm.  It was first explained by Matt Bishop, then a 2nd year graduate student at the University of Massachusetts.  Once the extrapolated data sets c/(c=0 and c/(q=0 are obtained, with appropriate uncertainties, lines are fit to them to give the zero angle intercept (related to inverse molecular weight) and slope (related to A2) and zero-concentration intercepts (again related inversely to molecular weight) and slope (related to Rg).  Of course, if c/( were plotted, we must reinsert the optical constant K first before useful parameters are extracted.  Also, these intercepts and slopes have associated uncertainties that translate into the final uncertainties of measurement.

Thus, from the NANG x NCONC data points, finally emerge just four parameters (with associated uncertainties):

- Bc = intercept of c=0 extrapolation (also Bc)	{17 a-d}

- B= intercept of =0 extrapolation (also B)

- Mc = slope of c=0 extrapolation (also Mc)

- M = slope of =0 extrapolation (also 

At this point, it’s only a short step to get Mw, Rg and A2.  Zimm.PAS can do this for you, but it’s a chancy game!  You are advised to take the “naked plot” (i.e., without fitted lines) and obtain the answers graphically as a check, if you are using Zimm.PAS.  The reason is that curved lines are difficult to recognize in Zimm.PAS, due partly to poor screen resolution and partly to the fact that data en grid sometimes tend to look linear when they are not.  Also, there are usually a few points that should be deleted, and this is easier to do by eye.  You can make several copies of the “naked grid” and draw your own lines several times to obtain the intercept and slopes and related uncertainties.

Once the four parameters and associated uncertainties are known, then it is simple to proceed based on Eqs. 5-6 (Kc/( at c=0 and q=0).

{18}

{19}

In this form, Mw is a function of the intercepts Bc and Band optical constant, K Thus,

{20}

Now, the only major uncertainty in the optical constant K comes from (dn/dc).  Thus,

{21}

Thus,

{22}

For manually drawn plots, one could simply report the difference between Mw from the c=0 and q=0 intercepts.  (However, for a least squares fit, there won’t be any difference!).

The radius of gyration may be calculated from:

Mc = slope of c=0 line = 		{23}

One could wonder what Mw is doing here.  An apparent radius can be obtained without knowing the molecular weight, as in the Guinier or Debye methods used in program GuiDe.VB.  But in Eq. 23, Mw appears because it is inversely proportional to Iq=0.  Clearly, if the graphically obtained intercepts differ, then the one from c=0 should be used in Eq. {23}.  So we write:

; Mc = slope of c=0 line

If K was evaluated in c.g.s. units (everything in polymer solution characterization should be!) then Rg comes out directly in cm.  The squared uncertainty in Rg is given by:

{24}

Finally, the virial coefficient is evaluated from MHowever, if q=0 then the slope reported is [d(c/()/dkc], where k is the scaling constant.  Correcting for the scaling, we have A2 = kKM/2.  If K was in c.g.s, then A2 is in mL-mol-g-2.  The uncertainty is:

In Summary:

Molecular weight:

Radius of gyration:

; Mc = slope of c=0 line

Virial coefficient:

A2 = kKM/2

Depolarized Corrections:  All the above is modified when the scatterers are optically anisotropic.  In this case, consult papers by Berry.

GuiDe:  It is important to remember that GuiDe doesn’t follow any of the above prescription at all!

Appendix 2.  GAL file structure

There follows the file PSBAD.GAL (aka run4157.gal)  The italicized color text does not appear in the real file; it’s just there to explain.

A:RUN4157.GAL 	The name of the file

11:23:22   		start time

03-18-1996		start date

18:14:14		end time

03-18-1996		end date

DOW 1683 POLYSTYRENE IN TOLUENE.............................first label

WIESLAW WORST DAY...........................................second label

WS........		operator

4880 			wavelength in Angstroms

TOLUENE.......	Rayleigh standard material

1.495 			refractive index of Rayleigh standard

1.495 			refractive index of solvent for the polymer solutions

1 			m = volume exponent (see documentation)

24.9 			lowest temperature during measurement

25.1 			highest temperature during measurement

1018.118 		dark count

24133.97 		Rayleigh standard intensity = Istd

120.0893 		uncertainty of Istd

90 			angle where Rayleigh standard was measured

9 			number of angles = Nang

9 			number of concentrations = Nconc; includes solvent as c=0

0 			c1

0.0001305 		c2

0.0002610 		c3

0.000654 		c4

0.001305 		c5

0.001957 		c6

0.002610 		c7

0.003914 		c8

0.005219 		c9

20 			first scattering angle ; next row is intensity and after that sintensity

81428.38  97591.7  122109.7  162909.5  197399.6  252660.4  301346.8  345541.3  373001.3

189.3396  1144.371  75.31445  854.0734  235.379  396.2553  522.5136  2446.77  1101.427

30

55574.58  66837.77  84346.55  112079.5  136769.3  177428.5  214188.9  232087  261054

197.434  392.5005  324.2591  173.9281  403.8692  568.4637  619.549  392.1764  669.6632

40

43341.26  51996.1  65401.65  87735.49  106350.9  138269.6  166658.1  179940.2  203643.3

239.8559  261.1551  328.7555  248.4714  564.8671  214.7371  180.5043  537.835  635.0795

50

36311.72  43244.43  55113.06  73145.96  89001.34  116999.8  137346.4  150733.3  170081.7

78.70634  234.4586  90.72663  191.5141  199.758  266.8599  273.3365  371.6528  367.2414

60

32257.47  38655.18  48411.66  62924.06  78320.7  102318.9  119509.9  132199.1  148991

106.8938  102.5643  131.3934  215.5929  239.1338  440.27  302.6808  93.53061  286.8959

75

28955.09  33923.45  42198.49  55891.61  68934.62  90674.13  105791.5  116571.6  131802.6

150.9799  122.7043  159.0369  253.9555  388.7278  110.6202  293.2361  286.2568  516.4243

90

27126.31  32055.37  39665.98  51456.23  62408.91  84488.31  98041.87  105803.9  122512.3

95.50737  194.1532  109.4389  230.5697  135.4062  152.2081  149.2928  281.5163  146.4149

110

29638.82  34479.8  42685.7  54635.88  67755.86  89103.03  104593  114505.7  130324.7

96.0753  157.568  165.4995  108.4058  114.0652  86.36407  222.5477  229.0522  273.5556

120

32521.76  36138.69  46225.73  59356.04  73573.95  95478.31  110844.5  122913.1  140521

170.4942  112.1322  216.0858  107.3936  102.3973  236.4535  258.8731  319.9878  372.4431

Appendix 3.  Preparing solutions worth measuring.

1.  Get clean cells!  Clean is relative and solvent-dependent.

For aqueous work, of not-very-crucial nature (latexes, colloids, other strong scatterers)

- soap the cells in hot water.

- use a gentle scrubber like a soft pipe cleaner.

- rinse in Nanopure water (from the ultrafilter, preferably).

- sonicate.

- rinse again.

- visualize in the scattering apparatus to see if clean (you should look at the whole length of the beam for many seconds and see no dust, at a low angle (like 45 degrees) and with a bright laser (argon ion or diode).

- pour the water out in HEPA chamber (so dusty air doesn’t get trapped).

- wrap in rinsed Al foil, again doing this in a clean environment.

- dry with heat or vacuum oven.

- load the cells in a clean environment.

- We don’t generally re-use ordinary cells after they have contained water-soluble polymers.

For aqueous work of a more critical nature (protein, dendrimers, dextran etc.)

- Consider silanation of the cells.

- DO NOT DO THIS WITHOUT CONSULTING WITH PROFESSOR RUSSO.  DISPOSAL OF THE SILANE SOLUTION NEEDS TO BE DONE CORRECTLY.

- It may be permissible to re-use silanated cells.  Silanated cells rinse clean and do not attract dust as quickly as others.

- You may still wish to work in the HEPA environment when possible.

For nonpolar organic solvents (cyclohexane, toluene, etc.)

- This is the easiest case.  Choices are:

wash in soap & water, rinse and place on acetone percolator

soak in chromerge, then rinse/sonicate/re-rinse and test/dry

- We don’t usually re-use such cells.  If for some reason, you wish to re-use a cell, make sure you wash it thoroughly in a good solvent for the polymer it once 	contained.

For polar organic solvents (DMF, THF, m-cresol, pyridine, etc.)

- Go wild!  use the chromerge & rinse, maybe taking care that some rinsing is done in the HEPA device.  Inspect the cells carefully.  Use extra sonication steps, maybe.  Pray.  Often.

2.  Be sure you can get solvent into the cells and still have it be clean  (As usual, Measure Nothing First).

For water, no problem--use Nanopure water.  You can check if the Nanopure system is working by putting some in a plastic cell and testing that in the apparatus (REMOVE TOLUENE VAT FIRST!).

For buffer solutions, it is much harder.  You may wish to de-dust simple salts by filtering concentrated solutions and evaporating to dryness.  This may leave waters of hydration, however.  You can rinse salts in filtered acetone to remove oils.  You can use very fine filters with some success.  Recycling the salt solution through a peristaltic or other pump has proven helpful sometimes--but you have to be careful not to pick up junk from the pump.  Very high-speed centrifugation (in Biochem ultracentrifuges, not our prep centrifuges) can be helpful.  Working in HEPA environment probably helpful.  It is hard to get buffers really clean.  To keep them clean, consider adding 3 mM sodium azide; bugs grow on almost any buffer.  Buffers can look clean and still misbehave!  Always test buffer solutions to see if they have any correlated DLS signal, even if you are planning only SLS experiments.  If you don’t know what this means, ask Professor Russo, who learned this lesson the hard way.

For nonpolar organics--ordinary PTFE filters are usually OK, the smaller the better, except that the Anotop filters (0.02 m) sometimes require pre-testing.  New 0.1 m PTFE filters from Whatman are the definite maximum size!

For polar organics--seriously consider distilling these solvents.  A vacuum distillation procedure appears in the Mark DeLong paper.  It is a good one.  Direct filtration works OK for THF if very small pore size (like 0.1 m or less) is used.  The big, pressurized vessel for the chromatography apparatus is helpful here, or the small Gelman vessels can be used.

3.  Making clean solutions at accurate concentrations.

So much for the easy part.  Polymers often carry lots of dust.  Because of their size, polymers prevent the use of very fine filters.  So, it is not always possible to take clean solvent, add it to a polymer, filter the solution and get a clean solution.  (There are cases where things work this simply, like polystyrene in toluene, but most systems pose greater difficulties).  It is often helpful to preclean the polymer.  To do this, dissolve it in clean solvent (it must be clean, even though polymer will add dust anyway; why?) and then precipitate it in clean non-solvent.  Wrap in rinsed aluminum foil or other clean container and vacuum dry.  Sometimes, precleaned polymer can be added to clean solvent and you will get a clean solution.  Yaaay!  More commonly, the solution will be gross, despite all your trouble.

There are limits to how much you can clean a dirty polymer solution.  As already mentioned, small filters impose the risk that polymer would be lost on the filter (actually, this is always a risk, due to adsorption, but small filters adds physical retention to the list of things that can go wrong during filtration; can you think of still others?).  Centrifugation sometimes works, but sometimes (esp. polar organics) the “dust” seems to have about the same density as your solution.  Also, you can alter the concentration by length centrifugation if your polymer density is very different from that of the solvent.  This is not usually a problem for polymer solutions in preparative cenrifuges, which provide all the g-field a light scattering cell can take.  Concentration change during centrifugation is definitely a potential problem if you are using one of the ultracentrifuges in Biochemistry--and you will also need something stronger than a light scattering cell in this case.  If you must withdraw solution after centrifugation, it is best to do so from the middle of the tube; some things sediment and some cream.  The good stuff is in the middle.  One thing that helps with dirty solutions is low-convective cells.  Try loading the cells only partially, using cells of smaller diameter, etc.  Dust that stands still--and not right in your viewed volume--is almost harmless.  But don’t give up easily; cleaner is always better.

The next issue is concentrations and accuracy thereof.

- Desired concentration range.  Lower is generally better...as long as you are accurate about concentrations and solvent subtraction.  For SLS, the solutions should scatter a few times more brightly than the solvent does. An advantage of a flow cell design over multiple inserted cells is that the stray light is constant.  Other things being equal, you can get to lower concentration in a flow cell.  Also consider making dilutions right in a single cell that is not moved during the whole measurement.  Slow and tedious, but possible.

- Volume or weight?  The light scattering equations are worked out for concentration as weight/volume--e.g., g/mL.  There are many ways to know this number, involving various levels of approximation.  A really great way to make your solutions would be to use about 1 liter of solvent to prepare the stock solution, make all subsequent dilutions with 25 or 50 mL precision pipettes and good (even calibrated!) volumetric glassware that has not been ruined in a thermal oven.  Well...that’s a lot of solvent, not to mention polymer!  We have had good luck doing the same thing on a smaller scale with Pipetman pipets.  It’s a very good idea to back up this volumetric scheme with mass measurements, because the weight precision is always better than volumetric measurement.  This basic protocol has worked OK.

- weigh 10-50 mg of precleaned polymer into precleaned 10-mL volumetric flask.  (Less for high-mass, strongly scattering polymers, more for very small or weakly scattering polymers).

- add half the solvent (CLEAN solvent, of course)

- allow to dissolve; this can take a long time, so be choosy about it!

- bring solvent up to the volumetric flask’s indicator line with more CLEAN solvent

- make sure the solution is totally homogeneous (after all traces of the swollen polymer are invisible, 25 tilts of a volumetric will do it.  Diffusion/convection will never do it).

- You now know Cmax.  If it’s clean, you can directly use it.  If not, you have to clean it totally (centrifuge, filter, etc.) and, probably, check to see if you have changed the concentration (by absorbance or by weighing a portion to dryness or checking whether the refractive index has changed during cleanup, using the differential refractometer).  Anyway, in this procedure you really must start with a clean Cmax stock solution.

- Goal is to make 0.8Cmax, 0.6Cmax, 0.4Cmax, 0.2Cmax, etc.  Here’s how.

- label each cell (WITH DIAMOND PENCIL OR GRAPHITE PENCIL NOT WITH INK)

- weigh each cell

- if you are dealing with a highly volatile solvent (THF, for example) you may

wish to label each cap and include them in the weighing also.

- if you use the caps, be sure they don’t get mixed up from one cell to the next; not all caps weigh the same.  Also be advised to work quickly; cap weight can change over time, which is why--if you have a sufficiently nonvolatile solvent--it is probably better to weigh quickly without the caps.  This is especially true if you might later want to return to the samples and modify their concentrations.

- rinse a pipet tip carefully with clean solvent and carefully empty

- add 0.8, 0.6, 0.4, 0.2 mL of stock to the weighed cells

- weigh each cell and the added stock solution

- add 0.2, 0.4, 0.6, 0.8 mL of clean stock

- weigh each cell again

- Now you have to calculate the concentrations of all the solutions.  You can do this by assuming volume additivity and the perfection of your pipet technique, not to mention the pipets themselves.  In this case you get 0.8Cmax, 0.6Cmax etc. directly.  You should probably check these concentrations using the weights, which are more precise than the volumes.  For example, to check the “0.8Cmax” solution compute Cmax ( (gstock - gempty cell)/(gfinal - gempty cell) where g represents mass in grams.  In using either of these methods, you are assuming that the partial specific volume of solvent and polymer do not change as a function of concentration in the measured range.  So if you really wish to be picky, you would know these values (from literature or densitometer measurement) and compute the concentrations from the masses in the usual way:  c2 = g2/(g1v1 + g2v2) where 2 represents polymer and 1 represents solvent.

9.  The Wyatt DAWN manual provides additional good information.

10.  Experience is the best, if a rather stern, teacher.

# Text boxes

c

q2

Kc/(

c1

c2

c3

c4

q2 + kc

Kc/(

c1

c2

c3

c4

q2 + kc

Kc/(

q12

q62

c1

c2

c3

c4

q2 + kc

Kc/(

c1

c2

c3

c4

q2 + kc

Kc/(

kc4

kc1

z

vertical

Laser

y



detector

x

beam

beam

m=2

m=1